## Slide 1
Particulate Matter
Sampling and Analysis
PM sampling, gravimetry and chemical analysis
Manos Manousakas
NCSR “Demokritos”, Greece
NextAIRE

## Slide 2
Scope of the presentation
OVERVIEW
This deck covers how to collect representative PM samples and how to convert them into reliable mass and chemical-composition data.
1
Sampling design
Choose PM fraction, locations, duration, filter media and field QA/QC based on the scientific or regulatory question.
2
Mass concentration
Measure PM mass using controlled filter conditioning, gravimetric weighing, flow checks and blank correction.
3
Chemical analysis
Analyze filters for elements, ions, carbonaceous material, organics and optical properties using complementary methods.
4
Data quality & mass closure
Validate the dataset, quantify uncertainties, identify artifacts and reconstruct measured PM mass from major components.
5
Reporting
Report data transparently: sampling conditions, analytical methods, detection limits, uncertainties and comparability.

## Slide 3
Why PM measurements matter
CONTEXT
Health
Inhalable particles reach the respiratory tract
Fine particles can penetrate deeply into the lungs
Composition matters: metals, carbonaceous material and organics can influence toxicity
Climate
Scattering and absorbing aerosols alter radiation
Particles act as cloud condensation or ice nuclei
Black carbon absorbs light and warms the atmosphere
Air-quality management
Mass concentration tells us how much PM is present
Chemical analysis tells us what it is made of
Reliable data underpin standards, trends and mitigation plans
Key point: sampling and analysis quality determines how much confidence we can place in PM trends, exceedances, health assessment and control measures.

## Slide 4
Current PM benchmarks used in air-quality assessment
CONTEXT
Reference and guideline values are becoming more stringent, which increases the need for robust sampling, calibration and uncertainty control.
Framework | PM2.5 annual | PM2.5 24-hour | PM10 annual | PM10 24-hour | Use
WHO AQG 2021 | 5 µg/m³ | 15 µg/m³ | 15 µg/m³ | 45 µg/m³ | Health-based guideline
EU revised directive | 10 µg/m³ by 2030 | — | 20 µg/m³ by 2030 | 45 µg/m³ | Legal limit values
US EPA NAAQS 2024 | 9.0 µg/m³ | 35 µg/m³ | — | 150 µg/m³ | National standards
For training purposes, compare measured concentrations with the framework relevant to your study area. Always check the applicable national transposition, averaging time and data-capture requirements.
The lower the benchmark, the more important it becomes to document blanks, flow uncertainty, weighing repeatability and chemical-method detection limits.
Sources: WHO AQG 2021; European Commission AAQD 2024; US EPA PM NAAQS 2024

## Slide 5
What are aerosols and particulate matter?
BASICS
Aerosols are liquid or solid particles suspended in a gas. In ambient air, we commonly refer to these particles as particulate matter (PM).
1
Dust
Solid particles formed by mechanical disintegration of parent material: soil, road dust, construction dust, mineral dust.
2
Smoke and soot
Combustion aerosols from fossil fuel, biomass burning and industrial processes. Often carbon-rich and light absorbing.
3
Mist and droplets
Liquid aerosol formed by condensation, atomization or hygroscopic growth of salts and inorganic aerosols.
Ambient PM is a mixture: primary emitted particles plus secondary particles formed in the atmosphere from gases such as SO₂, NOx, NH₃ and VOCs.

## Slide 6
PM is characterized by aerodynamic size
BASICS
Size determines where particles deposit, how they are sampled, and which chemical processes dominate.
0.001 µm
0.01
0.1
1
10
100 µm
Ultrafine / nuclei mode
Accumulation mode
Fine PM / PM2.5
Coarse PM
PM10
TSP
Aerodynamic diameter is not simply geometric size: it describes how a particle behaves in air and through size-selective inlets.

## Slide 7
Why size-selective sampling is essential
BASICS
Different size fractions represent different exposure pathways and sources, so PM10, PM2.5, PM1 and ultrafine-particle measurements answer different questions.
1
PM10
Inhalable fraction. Includes coarse particles such as crustal material, sea salt, road dust and some biological particles.
2
PM2.5
Fine particles. Often enriched in combustion products, secondary inorganic aerosol, organic aerosol and long-range transported material.
3
Ultrafine particles
Very small particles, often reported by number concentration rather than mass; relevant for fresh emissions and nucleation processes.
Select the size fraction before selecting the sampler, filter and analysis. A mismatch between scientific objective and inlet choice is difficult to correct later.

## Slide 8
Primary and secondary PM
BASICS
1
Primary PM
Particles emitted directly into the atmosphere. Examples: soot, mineral dust, sea spray, brake and tyre wear, fly ash, primary organic aerosol.
2
Secondary PM
Particles formed in the atmosphere from gas-phase precursors through chemical reactions and gas-to-particle conversion.
SO₂
NOx
NH₃
VOCs
oxidation + neutralization + condensation
SO₄²⁻
NO₃⁻
NH₄⁺
SOA
Analysis must preserve both primary and secondary components: volatile nitrate, ammonium and semi-volatile organic compounds can be affected by sampling conditions.

## Slide 9
What PM is made of
BASICS
A complete PM dataset usually requires multiple filters and/or multiple analytical techniques because no single method captures all components.
Secondary inorganics
SO₄²⁻, NO₃⁻, NH₄⁺
Carbonaceous matter
OC, EC, BC, organics
Sea salt
Na⁺, Cl⁻, Mg²⁺
Mineral dust
Al, Si, Ca, Ti, Fe
Trace metals
V, Ni, Cu, Zn, Pb, As, Cd
Water / other
particle-bound water, unmeasured mass
Practical implication: choose filter material according to target analytes, not only according to PM mass measurement.

## Slide 10
End-to-end workflow for PM sampling and analysis
SAMPLING
1
Define objective
Compliance, exposure, chemical speciation, trend, episode or research campaign
2
Design sampling
Site, PM fraction, duration, frequency, filters, instruments and blanks
3
Collect samples
Calibrated flow, documented metadata, controlled handling and storage
4
Analyze
Gravimetry, elements, ions, EC/OC, organics and optical properties
5
Validate & report
Uncertainty, mass closure, flags, limits, comparability and interpretation
Every step should be planned before the campaign begins. Analytical choices cannot fully compensate for poorly documented sampling.

## Slide 11
Designing a PM sampling strategy
SAMPLING
Define the question: regulatory compliance, trend monitoring, source indicators, exposure, method comparison or process study.
Select the PM fraction and averaging time before choosing the instrument.
Include representative site types: regional background, urban background, traffic, industrial, indoor or microenvironment as needed.
Plan temporal coverage: seasons, weekdays/weekends, episodes, wildfire/dust periods and meteorological variability.
Define QA/QC from the start: field blanks, travel blanks, duplicates, flow checks, weighing checks and storage conditions.
Minimum metadata for every sample: site, inlet/fraction, sampler ID, filter ID, start/end time, flow, pressure/temperature, RH, weather, local activities, operator, sample condition and storage.
Common mistake: collecting many samples with insufficient blanks and poor flow documentation. This can make the dataset difficult to defend later.

## Slide 12
Selecting the PM fraction and inlet
SAMPLING
Size-selective inlets define the measured quantity. The inlet and flow rate are part of the method, not just accessories.
Inlet cut-point
D₅₀ is the aerodynamic diameter at which 50% of particles are collected/transmitted by the size selector.
PM10 / PM2.5 / PM1
Fraction | Typical use | Important considerations
TSP | Bulk suspended particles | Weak size selectivity; historic datasets
PM10 | Regulatory inhalable fraction; coarse + fine PM | Crustal and sea-salt contributions often important
PM2.5 | Fine PM; health, combustion and secondary aerosol studies | Semi-volatile nitrate and organics need artifact control
PM1 / UF | Research, fresh emissions, aerosol dynamics | Specialized instruments and number-size metrics often needed
Reference-method details depend on the applicable standard, e.g., EN 12341:2023 for PM10/PM2.5 gravimetry.

## Slide 13
Sampler types and what they are best for
SAMPLING
Sampler type | Typical output | Strengths | Limitations
Low-volume sequential sampler | 24 h PM10/PM2.5 filters | Reference-style gravimetry; robust unattended operation | Limited mass for extensive chemistry
High-volume sampler | Large-area filters | High mass loading; useful for organics, trace components and multiple analyses | More handling; higher blank/artifact sensitivity
Dichotomous sampler | Fine + coarse fractions | Separates size fractions simultaneously | Lower flow/mass; maintenance of impactor/cut-point
Cascade impactor | Size-resolved samples | Particle-size distribution by mass/chemistry | Bounce, overloading and cut-point sensitivity
Online instruments | Minutes to hourly data | Events, diurnal patterns, rapid response | Calibration/equivalence and drift control needed
A good campaign often combines reference filter sampling with online measurements for temporal structure and offline methods for detailed chemistry.

## Slide 14
Typical filter-sampling train
SAMPLING
Size-selective inlet
Filter holder
Flow control
Pump
Data logger
The flow rate determines the inlet cut-point and sampled volume.
Leaks, filter loading and pump instability can bias concentration and particle size selection.
Meteorological sensors and pressure/temperature logging improve volume correction and data screening.
Filter holders and cassettes should minimize particle losses and contamination.
Field check: flow before and after sampling; leak check after filter installation; document alarms, power outages and abnormal noise/vibration.

## Slide 15
Choosing filter media
SAMPLING
Filter choice should be driven by the target analytes and the method requirements.
Filter medium | Best suited for | Advantages | Typical cautions
PTFE / Teflon | Gravimetry, XRF, ions | Low blank mass, chemically inert, good for non-destructive XRF | Static charge; fragile; not ideal for EC/OC
Quartz fiber | EC/OC, organic compounds, high-volume sampling | High temperature tolerance; common for carbonaceous aerosol | Carbon blank; adsorption artifacts; needs pre-baking
Nylon | Ions, especially nitrate | Good collection of some ionic species | May retain gases; not universal for metals
Polycarbonate | Microscopy, SEM/EDX, single particles | Smooth surface and defined pores | Limited loading; pressure drop
Glass/cellulose fiber | Mass/high-volume legacy applications | High capacity | Higher blanks; limited trace-element suitability
Use field blanks for each filter type. Filter blanks can dominate uncertainty for low-concentration sites or short sampling durations.

## Slide 16
Filter handling, conditioning and storage
SAMPLING
1
Before sampling
Unique filter ID, pre-conditioning, pre-weighing, lab blank record, clean transport cassette.
2
During sampling
Powder-free gloves/tweezers, avoid touching exposed surface, note damage or unusual loading.
3
After sampling
Condition before weighing, store at appropriate temperature, protect from light and contamination.
4
For chemistry
Split filters only with clean tools; store aliquots for ions/carbon/organics according to method needs.
Consistent conditioning is critical for mass: temperature/RH changes and hygroscopic water uptake can create apparent PM differences.

## Slide 17
Gravimetric determination of PM mass
ANALYSIS
Gravimetry is the reference basis for PM mass concentration: particles are collected on a filter and quantified by the mass increase.
mPM = (mloaded − munloaded) − blank correction
CPM = mPM / Vair
Use calibrated microbalance and anti-static control.
Condition filters before both pre- and post-weighing.
Use repeated weighing criteria and blank correction.
Report concentration in µg/m³ with volume basis specified.
EN 12341:2023 defines a standard gravimetric method for PM10/PM2.5 by sampling on filters and weighing them; the method includes specified inlet designs and nominal sampling conditions.
Source: EN 12341:2023 summary

## Slide 18
Air volume and concentration calculations
ANALYSIS
PM concentration depends on both measured mass and the sampled air volume. Flow uncertainty is often a major contributor.
Vair = Qavg × t
Qavg: average flow rate
t: valid sampling time
Vair: sampled volume at specified conditions
Record start and end time automatically where possible.
Check initial and final flow; flag deviations and instrument alarms.
State whether volumes are reported at ambient, standard or reference conditions.
Apply pressure and temperature corrections consistently when required.
Example: a 24 h sample at 2.3 m³/h collects about 55.2 m³ of air before any invalid-time corrections.

## Slide 19
QA/QC: what makes PM data defensible?
ANALYSIS
1
Blanks
Laboratory blanks, field blanks and travel blanks for every filter type and campaign phase.
2
Replicates
Duplicate or collocated samples to assess precision and sampling comparability.
3
Calibration
Flow calibrator, balance verification, instrument calibration standards and control charts.
4
Detection limits
Method detection limits based on blanks and realistic sample processing.
5
Uncertainty
Combine weighing, flow, blank, calibration and analytical uncertainty contributions.
6
Flags
Power failure, low flow, damaged filter, overload, precipitation, construction activity, contamination.
A QA/QC plan should be written before sampling starts and must be reflected in the data file, not only in the final report.

## Slide 20
Sampling artifacts and particle losses
SAMPLING
Artifacts are changes caused by the sampling method itself. They can affect both mass and chemistry.
1
Positive artifacts
Adsorption of organic vapors or reactive gases on filters; particle-bound water uptake.
2
Negative artifacts
Volatilization of nitrate, ammonium chloride and semi-volatile organic compounds during sampling or storage.
3
Particle losses
Inlet losses, tubing deposition, bounce in impactors, overloading, leaks and poor filter sealing.
4
Mitigation
Denuders, backup filters, appropriate face velocity, temperature control, clean handling and realistic blank correction.
Artifact control is especially important when measuring nitrate, ammonium, organic carbon and semi-volatile organic compounds.

## Slide 21
Field metadata: the invisible part of the sample
SAMPLING
The physical filter is only half of the sample. The metadata allow interpretation, QA screening and reproducibility.
Category | Examples to record
Site and sampler | Coordinates, site type, inlet height, sampler/inlet ID, filter holder ID
Time and flow | Start/end, duration, initial/final/average flow, flow alarms, total volume
Meteorology | Temperature, pressure, RH, precipitation, wind speed/direction where available
Sample condition | Filter ID, medium, lot, damage, visible deposits, overload, transport and storage conditions
Local context | Nearby activities, traffic changes, construction, fires, dust events, power interruptions
Good metadata turn samples into defensible environmental measurements. Poor metadata turn measurements into isolated numbers.

## Slide 22
Chemical analysis: matching methods to questions
ANALYSIS
Different techniques measure different fractions of PM chemistry. Combining methods produces a more complete picture.
Elements
XRF, PIXE, ICP-MS/OES, AAS
Al, Si, Fe, Ti, V, Ni, Cu, Zn, Pb, As, Cd
Ions
Ion chromatography
SO₄²⁻, NO₃⁻, NH₄⁺, Cl⁻, Na⁺, K⁺, Mg²⁺, Ca²⁺
Carbon
Thermal-optical EC/OC; optical BC
OC, EC, TC, equivalent black carbon
Organics
GC-MS, LC-MS, AMS/ACSM, EESI
Molecular markers, organic aerosol composition
Morphology
Microscopy, SEM/EDX
Size, shape and single-particle composition
Analytical plan = filter material + mass loading + target analytes + detection limits + available sample area + storage conditions.

## Slide 23
Elemental analysis of PM filters
ANALYSIS
Technique | Sample preparation | Strengths | Typical limitations
XRF | Usually non-destructive; filter analyzed directly | Multi-element; preserves sample; good for routine networks | Higher limits for light elements; matrix/loading corrections
PIXE | Non-destructive ion-beam method | Sensitive for many elements; small spots possible | Requires accelerator facility; beam-time constraints
ICP-MS / ICP-OES | Acid digestion/extraction | Very low detection limits; many trace metals | Destructive; digestion recovery and contamination control
AAS | Digestion/extraction | Targeted metal analysis; widely available | Element-by-element; lower throughput
Elemental datasets are useful for crustal material, sea salt, industrial tracers, traffic-related metals and quality checks such as elemental ratios.
Source examples: EPA IO-3.3 for XRF; EPA IO-3.4 for ICP methods

## Slide 24
Water-soluble ions by ion chromatography
ANALYSIS
Ion chromatography quantifies the major water-soluble inorganic aerosol components after filter extraction.
Common target ions
SO₄²⁻
NO₃⁻
NH₄⁺
Cl⁻
Na⁺
K⁺
Mg²⁺
Ca²⁺
Use clean extraction vessels and known extraction volume.
Report extraction efficiency and blank levels where relevant.
Nitrate and ammonium are sensitive to sampling/storage artifacts.
Ion balance can reveal data issues or unmeasured species.
Secondary inorganic aerosol is often represented by sulfate, nitrate and ammonium; sea-salt and crustal contributions can be assessed through Na⁺, Cl⁻, Mg²⁺, Ca²⁺ and related elements.

## Slide 25
Carbonaceous aerosol: OC, EC and BC
ANALYSIS
Carbonaceous material is a major component of fine PM, especially in combustion-influenced environments.
1
Organic carbon (OC)
Carbon in organic compounds. Includes primary organics and secondary organic aerosol. Reported as carbon mass, not total organic mass.
2
Elemental carbon (EC)
Refractory carbon measured by thermal-optical methods. Often associated with combustion and diesel/soot-like particles.
3
Black carbon (BC/eBC)
Optically absorbing carbon measured by light absorption instruments. Method-dependent; not identical to EC.
Organic mass ≈ OC × OM/OC factor
Thermal-optical analysis requires quartz filters and controls for pyrolyzed carbon using optical transmittance or reflectance. Field blanks are essential.
Source example: NIOSH Method 5040 for elemental carbon / thermal-optical analysis

## Slide 26
Organic aerosol and molecular-level analysis
ANALYSIS
Organic PM is chemically complex. The appropriate technique depends on whether the objective is bulk carbon, molecular markers or real-time organic aerosol composition.
Approach | Typical information | Notes
GC-MS / LC-MS | Molecular markers, PAHs, hopanes, levoglucosan, polar organics | Requires extraction; method and standard availability matter
AMS / ACSM | Non-refractory submicron aerosol composition and mass spectra | Online/high-time-resolution; usually PM1 non-refractory fraction
EESI / soft ionization | Less-fragmented organic molecular fingerprints | Advanced technique for high-time-resolution organic aerosol composition
FTIR / NMR | Functional groups and bulk organic composition | Useful for complementary characterization
Storage matters: organic compounds can degrade, volatilize or be contaminated. Use clean procedures, low temperatures and light protection when required.

## Slide 27
Optical analysis: black carbon and light absorption
ANALYSIS
1
Aethalometer / MAAP / PSAP
Measure light attenuation or absorption on a filter/tape to estimate equivalent black carbon and absorption coefficients.
2
Advantages
High-time-resolution data, strong combustion signal, useful for traffic and biomass-burning indicators when combined with other measurements.
3
Cautions
Requires correction for loading, multiple scattering and wavelength dependence. BC is operationally defined and not interchangeable with EC.
4
Complementarity
Pair optical BC with EC/OC, XRF and ion data to understand carbonaceous and combustion-related PM more completely.
Report instrument type, wavelength(s), correction algorithm and assumed mass absorption cross-section when reporting eBC.

## Slide 28
Online and high-time-resolution PM composition
ANALYSIS
Online instruments complement filter sampling by resolving short episodes, diurnal patterns and rapidly changing atmospheric processes.
1
Xact / online XRF
Hourly or sub-hourly elemental concentrations
2
ACSM / AMS
Organics, sulfate, nitrate, ammonium and chloride in non-refractory fine PM
3
Aethalometer
Equivalent black carbon and absorption at multiple wavelengths
4
TEOM / BAM / OPC
Continuous PM mass or particle-number/size information
For regulatory use, automated instruments typically require equivalence demonstration, regular calibration/maintenance and comparison with reference measurements.

## Slide 29
Sample preparation: keep the analyte in mind
ANALYSIS
Preparation should extract the target species quantitatively while minimizing contamination, loss and chemical transformation.
Target | Typical preparation | Key controls
Mass | No extraction; conditioning and weighing | Humidity, static, balance drift, repeated weighing
Elements by XRF/PIXE | Direct filter analysis | Homogeneity, filter area, thickness/loading, calibration standards
Metals by ICP/AAS | Acid digestion or leaching | Acid purity, digestion recovery, vessel blanks, reference material
Water-soluble ions | Water extraction + filtration | Extraction volume/time, blank ions, storage temperature
EC/OC | Filter punch from quartz filter | Pre-baking, punch area, carbon blanks, split protocol
Organics | Solvent extraction or online inlet | Volatilization, degradation, matrix recovery, internal standards

## Slide 30
Data processing and validation
DATA
1
Blank correction
Subtract appropriate blank values and propagate blank variability to uncertainty.
2
LOD / LOQ
Define detection and quantification limits from method blanks or replicate low-level measurements.
3
Uncertainty
Document mass, volume, calibration, extraction, repeatability and blank uncertainty.
4
Consistency checks
Ion balance, elemental ratios, time-series continuity, collocated precision and mass closure.
5
Data flags
Use transparent flags for invalid samples, suspected contamination, under-load/overload and instrument issues.
6
Traceability
Keep raw files, lab notebooks, calibration certificates, sample chain of custody and processing scripts.

## Slide 31
Chemical reconstruction and mass closure
DATA
Chemical reconstruction checks whether measured components explain gravimetric PM mass and helps identify missing or uncertain fractions.
Reconstructed PM ≈ secondary inorganic aerosol + organic mass + elemental carbon + sea salt + mineral dust + trace elements + water/other
Secondary inorganic aerosol
SO₄²⁻ + NO₃⁻ + NH₄⁺
Organic mass
OC × OM/OC factor
Mineral dust
oxide-based formula or element-based estimate
Sea salt
Na⁺/Cl⁻/Mg²⁺ based estimate
Carbon
EC and/or BC as method-specific carbon fraction
Unmeasured
particle-bound water and other residual mass
Use chemical reconstruction as a validation and interpretation tool, not as a replacement for well-designed sampling and analytical QA/QC.

## Slide 32
Chemical reconstruction: assumptions and pitfalls
DATA
OM/OC conversion factors vary by location, source mix, oxidation state and season.
Estimating sulfate from total sulfur assumes sulfur is present as sulfate and fully particle-bound.
Sea-salt chloride can be depleted by reaction with nitric or sulfuric acid.
Mineral-dust formulas assume crustal oxides and may be biased by industrial or traffic metals.
Water associated with hygroscopic particles is usually not fully captured by offline chemical analysis.
A good mass closure does not guarantee each component estimate is unique or physically perfect.
Practical use: compare reconstructed mass with gravimetric mass, examine residuals, and investigate systematic differences by season, RH, PM fraction and site type.
Report formulas and factors explicitly. “Chemical reconstruction” is not a single universal equation; it is a family of approximations tailored to the dataset.

## Slide 33
Reporting PM sampling and analysis results
DATA
A complete report enables another laboratory or reviewer to understand exactly what was measured and how reliable the results are.
1
Methods
Sampler/inlet, flow, filter medium, sampling duration, standards and analytical techniques.
2
Results
Mass concentrations, chemical concentrations, uncertainty, LOD/LOQ and invalid/flagged samples.
3
Quality
Blanks, duplicates, calibrations, recovery checks, collocation statistics and data completeness.
4
Interpretation
Temporal patterns, mass closure, component fractions, episodes and comparison with benchmarks.
Use consistent units: µg/m³ for PM mass and components, ng/m³ for trace species where appropriate, and specify the air-volume reference conditions.

## Slide 34
Practical checklist for a PM campaign
CHECKLIST
1
Before fieldwork
Objective, site permissions, sampler maintenance, flow calibration, filter conditioning, blanks, labels, chain-of-custody forms.
2
During fieldwork
Flow checks, leak checks, inlet cleaning, meteorology, field logs, sample transport, field blanks, equipment alarms.
3
Laboratory
Conditioning, weighing criteria, extraction/digestion, calibration curves, reference materials, blanks, replicate analysis.
4
Data processing
Volume correction, blank correction, uncertainty, LOD/LOQ, flags, mass closure, final metadata table.

## Slide 35
Take-home messages
SUMMARY
PM sampling is a measurement system: inlet, flow, filter, conditioning, handling and metadata all define the result.
Filter media must match the analytical objective; one filter type rarely serves every purpose.
Gravimetric PM mass is simple in concept but sensitive to blanks, humidity, static charge and flow accuracy.
Chemical analysis requires method-specific QA/QC and careful sample preparation.
Chemical reconstruction and mass closure are powerful tools for checking data consistency and identifying missing fractions.
Transparent reporting of methods, limits, uncertainties and flags is essential for comparability and policy relevance.
Final principle: the strongest PM datasets are built at the planning stage, not after the samples are collected.

## Slide 36
Selected references and source material
REFERENCES
Original training material by notes from Manos Manousakas
World Health Organization, Global Air Quality Guidelines for PM2.5, PM10, O3, NO2, SO2 and CO, 2021.
European Commission, revised Ambient Air Quality Directive / cleaner air by 2030, 2024.
US Environmental Protection Agency, National Ambient Air Quality Standards for Particulate Matter, 2024 revision.
EN 12341:2023, Ambient air — standard gravimetric method for PM10 and PM2.5 mass concentration.
US EPA Compendium Methods IO-2.1, IO-3.3 and IO-3.4 for particulate sampling and elemental analysis.
NIOSH Method 5040 for elemental carbon / thermal-optical analysis.